## Slide 1: Approach of a UK auditor
Paul Mudway
Mudway Health, Safety & Environment
Notes: <number>

## Slide 2: Approach of a UK auditor
Auditing the stack monitoring team – key points
Manual methods
Instrumental methods
Proficiency testing schemes
Notes: <number>

## Slide 3: Auditing the stack monitoring team – key points
In the UK a typical UKAS audit is 2 days: 1 day site witness, 1 day laboratory
Purpose of the audit is to assess conformity with standards, documented procedures and guidance notes
In Ireland, relevant documents are ISO/IEC 17025, TS 15675, EN 15259, AG1 and AG2
Notes: <number>

## Slide 4: Auditing the stack monitoring team – key points (cont.)
Site-specific safety induction
Review risk assessment first (section 6 of AG1)
Permanent/temporary platform inspection certificate (HSA GA3)
Correct PPE being worn?
Site review (Appendix 4 of AG2 and EN 15259)
Work file (Appendix 6 of AG2)
Notes: <number>

## Slide 5: Auditing the stack monitoring team – key points (cont.)
Notes: <number>

## Slide 6: Auditing the stack monitoring team – key points (cont.)
Site-specific protocol (Appendix 5 of AG2 and EN 15259)
Documented procedures available?
Order of hierarchy? (CEN – ISO – National)
Copies of relevant sampling standards available (e.g. IE EN 13284-1)?
Notes: <number>

## Slide 7: Manual methods
Notes: <number>

## Slide 8: Manual methods (cont.)
Are inclined gauges zeroed and level?
Notes: <number>

## Slide 9: Manual methods (cont.)
Nozzle condition and size (≥ 6mm diameter for IE EN 13284-1 tests
Notes: <number>

## Slide 10: Manual methods (cont.)
Calibration status of all sampling train components
Pitot tube leak check (AG2 section 5.9)
Gas velocity traverse and swirl check?
Balance checked before weighing impingers?
Notes: <number>

## Slide 11: Manual methods (cont.)
Notes: <number>

## Slide 12: Manual methods (cont.)
Sampling train leak check: ≤ 2% of normal sampling flow rate (typically 2% of 1m3.h-1 = 0.33l.m-1 at vacuum (-10 to -15”Hg) 5% for PCDD/PCDF and PAH
Before and after test, if components are changed during test and when changing sampling lines for PCDD/PCDF tests
Leak checks also required for field blanks
Notes: <number>

## Slide 13: Manual methods (cont.)
Notes: <number>

## Slide 14: Manual methods (cont.)
Impinger/resin trap/sorbent tube inlet temperatures (refer to log sheets)
≤20°C for PCDD/PCDF
≤30°C for metals (can be useful!)
≤40°C for VOC sorbent tubes
Notes: <number>

## Slide 15: Manual methods (cont.)
Notes: <number>

## Slide 16: Data verification
Notes: <number>

## Slide 17
Where:
constant = 8.038 x 10-5
Cp= Pitot tube coefficient
Dn = nozzle diameter
M = molecular weight
Bw = moisture fraction (0 – 1.0 volume basis)
T = absolute temperature, K
P = absolute pressure (mmHg)
s = stack
m = meter
Data verification (cont.)
Notes: <number>

## Slide 18: Data verification (cont.)
Mean isokinetic flow rate during test:
-5% to +15% of ideal for EN 13284-1
±10% of ideal for ISO 9096
Notes: <number>

## Slide 19: Instrumental methods
Notes: <number>

## Slide 20: Instrumental methods (cont.)
Calibration status of analysers
Ambient temperature (e.g. 5°C to 40°C for Horiba PG-250)
Heated line temperature (160°C - 180°C)
Use of chiller dryer for NOx measurements using SRM (10% NO2)
NOx converter efficiency (alphanumeric PG-250)
Notes: <number>

## Slide 21: Instrumental methods (cont.)
Gas conditioner temperature (≤4°C)
ELV multiplier rule (1.5 x for WID, 2.5 x for LCPD) certified to EN 15267-3
Calibration sequence, before test zero – span – zero - sampling system verification (within 2% of span for CO, NOx, O2, 5% of span for SO2)
After test, sampling system verification only
Notes: <number>

## Slide 22: Instrumental methods (cont.)
Traceability of calibration gases
Check expiry dates
Notes: <number>

## Slide 23: Instrumental methods (cont.)
Notes: <number>

## Slide 24: Instrumental methods (cont.)
However, it is not necessary to use certified gases on site, provided that standards are traceable to the SI by means of an unbroken chain of calibrations or comparisons linking them to relevant primary standards of the SI units of measurement
Span gas concentrations should be approximately the half-hourly ELV or 50% to 90% of the analyser range
EN 12619 specifies 16 mg.m-3 propane span gas
Notes: <number>

## Slide 25: Instrumental methods (cont.)
Uncertainty quoted on certificate of analysis should be ≤±2%
Analyser drift over measuring period. ≤2% of span value is acceptable, 2% - 5% apply drift correction >5% reject results
Notes: <number>

## Slide 26: Data verification
t is the Celsius temperature;
p is the difference between the static pressure of the sample gas and the standard pressure;
h is the absolute water vapour content (by volume);
o is the oxygen content in dry gas (by volume);
os is the oxygen standard condition.
Notes: <number>

## Slide 27: Proficiency testing (PT) schemes
Participation is requirement of ISO/IEC 17025 accreditation
NPL particulate stack emissions
NPL calibration gas (cylinders)
NPL gas measurement (stack simulator)
www.npl.co.uk
Notes: <number>

## Slide 28: Proficiency testing (PT) schemes (cont.)
VSL Delft, Netherlands calibration gas (cylinders)
www.vsl.nl
All schemes are run twice per year
Notes: <number>